## Information
Information
Skriptversion | v45.1.0
Användare | jesper.fredriksson
Datum Tid | 2025-12-28 18:12
Minitab Macro | %D12547-MTBU-SWE
CSV INFO
CSV | 2099 09 09 GXMTB RIF ULTRA 92303 Tests Summary.csv
LSP | 2 (92302 x2, 92303 x207)
Antal tester | 209
Dublett Sample ID | 1 (C15NOV23D_02_1_14X x2)
Dublett Cartridge S/N | 2 (614386401 x2, 614332989 x2)
Använda INF/GX | GX1; GX3; GX5; Infinity-I; Infinity-V; Infinity-VI; Infinity-VIII
Bag Numbers Tested Using Infinity: | #00-10 x183
HEADER INFO
Worksheet | 2099 09 09 GXMTB RIF ULTRA 92303 Worksheet D25862.xlsx
Part No. | 700-5702 | Avvikande flik: Data - First 10
Batch No(s) | 1001506046 | Avvikande flik: Data - First 10
Cartridge No. (LSP) | 92303 | Avvikande flik: Data - First 10
Document Number | D25862 Attachment 1
Rev | AW
Effective | 9/11/2025
Seal Test POS | 2099 09 09 GXMTB RIF ULTRA 92303 Seal Test Pos.xlsx
Document Number | D10552
Rev | AU
Effective | 07/02/2025
Seal Test NEG | 2099 09 09 GXMTB RIF ULTRA 92303 Seal Test Neg.xlsx
Document Number | D10552
Rev | AU
Effective | 07/02/2025
LÄNKAR
SharePoint Batch | Öppna 1001506046
IPT App | LÄNK
MES Login | LÄNK
CSV Uploader | LÄNK
BMRAM | LÄNK
Agile | LÄNK
Note on A1: Generell spårbarhetsinfo samt Minitab-makro om tillgängligt.
Note on A7: Kort summering av vald .csv-fil.
Note on B14: Visar alltid Bag Number och summerar antal tester på Infinity.
Note on A16: Extreherar information från respektive fil och reagerar på avvikande innehåll.
Note on A33: Direktlänkar till funnet "Batch Number" i SharePoint (om valt) samt diverse IPT-relaterade verktyg.

## Seal Test Info
Seal Test Information
| Neg | Pos
ROBAL | ROBAL-006 | ROBAL-006
Part Number | 700-5702 | 700-5702
Batch Number | 1001506040 | 1001506046 | Mismatch
Cartridge Number (LSP) | 92302 | 92303 | Mismatch
PO Number | 1110580 | 1110580
Assay Family | MTB | MTB
Weight Loss Spec | 100 | 100
Balance ID Number | SCLSYS01, 003070 | SCLSYS03, SCLSYS04
Balance Cal Due Date | Apr-26 | Apr-26
Vacuum Oven ID Number | 003478, 003298 | 003346, 003587
Vacuum Oven Cal Due Date | Jun-26 | Jun-26
Timer ID Number | Nr. 18, Nr. 19 | Nr. 20, Nr. 21
Timer Cal Due Date | Mar-26 | Mar-26
Name of Tester
| Bincenth Salazar | Bincenth Salazar
| David Björklund | Katarina Ilic
| Evangelos Samakidis | Laith Ukra
| Mahdi Hussein | Leila Hussein
| Marina Ahmadi | Marina Ahmadi
| Warda Yousaf | Toni Shabo
Print Full Name, Sign, and Date
| Neg | Pos
| Jesper, FEST, 20990505 | Jesper, FEST, 20990505

## STF Sum
STF | Datasheet | Cartridge ID | Initial Weight (mg) | Final Weight (mg) | Weight Loss (mg) | Weight Loss Pass/Fail | Observation
NEG | Datasheet (1) | 01_0_08 | 37055.7 | 33197.2 | 3858.5 | FAIL | Lid Weld (Leak)
NEG | Datasheet (1) | 02_0_09 | 37438.3 | 37037.6 | 400.7 | FAIL | Broken Chimney
NEG | Datasheet (4) | 07_0_03 | 37645.1 | 37650 | -4.9 | Minusvärde | NA
NEG | Datasheet (4) | 08_0_04 | 37694.2 | 37393.7 | 300.5 | FAIL | Broken Chimney
POS | Datasheet (1) | 01_1_15 | 37112.7 | 37012.6 | 100.1 | FAIL | Film Seal (Leak)
POS | Datasheet (1) | 02_2_20 | 37472.4 | 37476.1 | -3.7 | Minusvärde | NA
POS | Datasheet (2) | 04_1_14 | 37668.5 | 37518.5 | 150 | FAIL | Lid Weld (Leak)
POS | Datasheet (3) | 05_1_15 | 37442.1 | 37341.1 | 101 | FAIL | Film Seal (Leak)
POS | Datasheet (5) | 10_1_13 | 37760.6 | 37160.2 | 600.4 | FAIL | Broken Chimney

## Utrustninglista IPT (hidden)
In Process Testing | [5 empty] | Kontroll och kalibrering
Senast uppdaterad: | [5 empty] | Utförd av Cepheid |  |  |  | Utförd av externt företag
Utrusningens namn: | Asset Type: | Asset ID/BMRAM ID: | Modell/serienr. | Serial # |  |  | Rum: | Nästa kontroll | Frekvens | Nästa kontroll | Frekvens | Frekvens | Nästa kontroll-datum
GeneXpert / Infinity
GeneXpert XVI - 16 Sites (Dell Computer), GX1 | GX16 Instrument | 000199 | GX16 Instrument | 802598 | N/A | N/A | Positivt labb 5.575 | 2026-03-26 | 2 gång/år | N/A | N/A
GeneXpert XVI - 16 Sites (Dell Computer), GX2 | GX16 Instrument | 000368 | GX16 Instrument | 802624 | N/A | N/A | Positivt labb 5.575 | 2026-03-26 | 2 gång/år | N/A | N/A
GeneXpert XVI - 16 Sites (Dell Computer), GX3 | GX16 Instrument | 000369 | GX16 Instrument | 802685 | N/A | N/A | Positivt labb 5.575 | 2026-03-26 | 2 gång/år | N/A | N/A
GeneXpert XVI - 16 Sites (Dell Computor), GX-7 | GX16 Instrument | 003474 | GX16 Instrument | 110012274 | N/A | N/A | Positivt labb 5.575 | 2026-03-26 | 2 gång/år | N/A | N/A
Infinity-I. | Genexpert Infinity-80 | 000375 | Infinity-80 | 802069 | N/A | N/A | Positivt labb 5.575 | Mar-25 | 2 gång/år | N/A | N/A
Infinity-I | Genexpert Infinity-80 | 000375 | Infinity-80 | 802069 | N/A | N/A | Positivt labb 5.575 | 2026-06-01 | 2 gång/år | N/A | N/A
Infinity-III | Genexpert Infinity-80 | 001163 | Infinity-80 | 807363 | N/A | N/A | Positivt labb 5.575 | 2026-04-26 | 2 gång/år | N/A | N/A
Infinity-V | Genexpert Infinity-80 | 002593 | Genexpert Infinity-80 | 839032 | N/A | N/A | Positivt labb 5.575 | 2026-01-01 | 2 gång/år | N/A | N/A
Infinity-VI | Genexpert Infinity-80 | 002796 | Genexpert Infinity-80 | 847922 | N/A | N/A | Negativt labb 5.409 | 2026-05-25 | 2 gång/år | N/A | N/A
Infinity-VIII | Genexpert Infinity-80 | 003066 | GeneXpert Infinity-80, PN 900-0503R | 803094 | N/A | N/A | Negativt labb 5.409 | 2026-01-01 | 2 gång/år | N/A | N/A
GeneXpert XVI - 16 Sites (Dell Computor), GX5 | GX16 Instrument | 003081 | GX16 Instrument | 110012284 | N/A | N/A | Negativt labb 5.409 | 2026-04-25 | 2 gång/år | N/A | N/A
GeneXpert XVI - 16 Sites (Dell Computor), GX6 | GX16 Instrument | 003082 | GX16 Instrument | 110012293 | N/A | N/A | Negativt labb 5.409 | 2026-04-25 | 2 gång/år | N/A | N/A
Sealtest
003478 | Vakuumugn | 003478 | VD56 + pump S/N 234346 | 20230000019821 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | 2026-06-01 | 1 gg/år
003298 | Vakuumugn | 003298 | VD56 + pump S/N 100603504-100556308 | 20230000003912 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | 2026-06-01 | 1 gg/år
003587 | Vakuumugn | 003587 | VD56 + pump S/N 233829 | 13-05094 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2026-06-01 | 1 gg/år
003346 | Vakuumugn | 003346 | VD56 + pump S/N 234348 | 20230000019820 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2026-06-01 | 1 gg/år
003070 | Våg | 003070 | XSR104 | C214830934 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | 2026-04-01 | 1 gg/år
SCLSYS01 | Våg | SWE-SCLSYS01-SCALE01 | XS104 / 850-0305 | B838450098 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | 2026-04-01 | 1 gg/år
SCLSYS03 | Våg | SWE-SCLSYS03-SCALE01 | XS104 / 850-0627 | B838450099 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2026-04-01 | 1 gg/år
SCLSYS04 | Våg | SWE-SCLSYS04-SCALE01 | XS104 / 850-0627 | B838450100 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2026-04-01 | 1 gg/år
Timer Nr. 18 | Timer | 003753 | EU 609-0206 | E2423931 |  |  | Negativt labb 5.409 | Mar-26 | Timers byts ut efter att kalibreringsdatum utgått | N/A | N/A
Timer Nr. 19 | Timer | 003754 | EU 609-0206 | E2423935 |  |  | Negativt labb 5.409 | Mar-26 |  | N/A | N/A
Timer Nr. 20 | Timer | 003755 | EU 609-0206 | D2421658 |  |  | Positivt lab 5.576 | Mar-26 | Timers byts ut efter att kalibreringsdatum utgått | N/A | N/A
Timer Nr. 21 | Timer | 003756 | EU 609-0206 | D2421660 |  |  | Positivt lab 5.576 | Mar-26 |  | N/A | N/A
Pippeter | [11 empty] | 1 gg/år | Apr-24
Multipette® nr. 13 | Multipette | 001361 | Multipette stream E3 | G42574F | N/A | N/A | Positivt lab 5.576 | 2026-03-26 | 1gång/3mån | N/A | N/A
Multipette® nr. 16 | Multipette | 002026 | Multipette E3 | M45826G | N/A | N/A | Positivt lab 5.576 | Jul-24 | 1gång/3mån | N/A | N/A
Multipette® nr. 17 | Multipette | 002077 | Multipette E3 | M37056H | N/A | N/A | Positivt lab 5.576 | 2026-01-31 | 1gång/3mån | N/A | N/A
Multipette® nr. 18 | Multipette | 002112 | Eppendorf Multipette ® E3x | L50185H | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Multipette® nr. 19 | Multipette | 002114 | Eppendorf Multipette ® E3x | L50216H | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Multipette® nr. 20 | Multipette | 002538 | Multipette E3 | Q21180I | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Multipette® nr. 21 | Multipette | 002539 | Multipette E3 | Q21012I | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Multipette® nr. 22 | Multipette | 003064 | Eppendorf Multipette E3 | G16044L | N/A | N/A | Positivt lab 5.576 | 2026-01-26 | 1gång/3mån | N/A | N/A
Multipette® nr. 34 | Multipette | 003077 | Multipette stream E3 | J17233L | N/A | N/A | Positivt lab 5.576 | 2026-02-25 | 1gång/3mån | N/A | N/A
Multipette® nr. 37 | Multipette | 002391 | Eppendorf Multipette E3 | G30024H | N/A | N/A | Positivt lab 5.576 | Sep-25 | 1gång/3mån | N/A | N/A
Multipette® nr. 31 | Multipette | 003056 | Eppendorf Multipette E3 | K46056K | N/A | N/A | Negativt labb 5.409 | 2026-01-26 | 1gång/3mån | N/A | N/A
Multipette® nr. 32 | Multipette | 003057 | Eppendorf Multipette E3 | K46102K | N/A | N/A | Negativt labb 5.409 | 2026-01-26 | 1gång/3mån | N/A | N/A
Multipette® nr. 33 | Multipette | 003065 | Eppendorf Multipette E3 | G16092L | N/A | N/A | Negativt labb 5.409 | 2026-01-26 | 1gång/3mån | N/A | N/A
Multipette® nr. 35 | Multipette | 003078 | Multipette stream E3 | J17307L | N/A | N/A | Negativt labb 5.409 | 2026-02-26 | 1gång/3mån | N/A | N/A
Multipette® nr. 36 | Multipette | 003079 | Multipette stream E3 | J17321L | N/A | N/A | Negativt labb 5.409 | 2026-02-25 | 1gång/3mån | N/A | N/A
Finnpipette F1 Nr.3 20-200µl | Finnpipette | 000466 | F1 20-200uL nr 3 | HJ04851 | N/A | N/A | Positivt lab 5.576 | 2026-01-26 | 1gång/3mån | N/A | N/A
Finnpipette F3 Nr 4 100-1000µl | Finnpipette | 002393 | 100-1000 ul Nr 4 | QH72852 | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Finnpipette Nr 6 20-200µl | Finnpipette | 002666 | 20-200 ul | K45707I | N/A | N/A | Negativt labb 5.409 | 2026-02-25 | 1gång/3mån | N/A | N/A
Finnpipette F3 Nr 7 100-1000µl | Finnpipette | 003126 | Research plus 1-channel | L19743I | N/A | N/A | Negativt labb 5.409 | 2026-02-25 | 1gång/3mån | N/A | N/A
Finnpipette F3 Nr 23 100-1000µl | Finnpipette | 003007 | Finnpipette F3 100-1000 ul | SH92611 | N/A | N/A | Negativt labb 5.409 | 2026-03-26 | 1gång/3mån | N/A | N/A
Finnpipette F3 Nr 24 100-1000µl | Finnpipette | 003006 | Finnpipette F3 100-1000 ul | SH92610 | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Finnpipette Nr 25 20-200µl | Finnpipette | 002109 | Eppendorf Research® Plus 1-channel | Q42723G | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Finnpipette Nr 26 2-20 µl | Finnpipette | 002111 | Eppendorf Research® Plus 1-channel | Q45392G | N/A | N/A | Positivt lab 5.576 | 2026-03-01 | 1gång/3mån | N/A | N/A
Övrigt
Referensvikt 1 g 003229 | Vikt | 003229 | 158397 - 1g OIML Class: E2 | C123934096 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | 2025-02-01 | 1 gg/år
Referensvikt 100 g (E2)003246 | Vikt | 003246 | ZG6L (E2) | S2301067 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | 2025-02-01 | 1 gg/år
Referensvikt 100g 000519 | Vikt | 000519 | NA | 12828038 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2025-02-01 | 1 gg/år
Referensvikt 1g 002937 | Vikt | 002937 | Item number 158397 | C123934097 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2025-02-01 | 1 gg/år
LAF 1 | LAF-bänk | 003055 | ScanLaf Mars 1200 Labogen Class 2 cabinet | 2224686 | N/A | N/A | Negativt labb 5.409 | N/A | N/A |  | 1 gg/år
LAF 2 | LAF-bänk | 003059 | ScanLaf Mars 1200 Labogen Class 2 cabinet | 2224687 | N/A | N/A | Negativt labb 5.409 | N/A | N/A |  | 1 gg/år
LAF 3 | LAF-bänk | 003060 | ScanLaf Mars 1200 Labogen Class 2 cabinet | 2224688 | N/A | N/A | Negativt labb 5.409 | N/A | N/A |  | 1 gg/år
LAF 3 | LAF-bänk | 000412 | Mars 1200 | L.04110289 | N/A | N/A | Positivt lab 5.576 | N/A | N/A |  | 1 gg/år
LAF 5 | LAF-bänk | 001747 | Mars Premium 1200 | L.11150043 | N/A | N/A | Positivt lab 5.576 | N/A | N/A |  | 1 gg/år
LAF 4 | LAF-bänk | 002441 | Cabinet Biosafety Biovanguard | G19118701-01 | N/A | N/A | Positivt lab 5.576 | N/A | N/A |  | 1 gg/år
LAF 2 | LAF-bänk | 002897 | Esco | 2015-99627 | N/A | N/A | Positivt lab 5.576 | N/A | N/A |  | 1 gg/år
FQC LAF | LAF-bänk | 002898 | ESCO | 2015-99629 | N/A | N/A | Positivt lab 5.576 | N/A | N/A |  | 1 gg/år
003407 | VBA Venting Fixture | 003407 | 700-8285 | N/A | N/A | N/A | Positivt lab 5.576 | N/A | N/A | N/A | N/A
003406 | VBA Venting Fixture | 003406 | 700-8285 | N/A | N/A | N/A | Negativt labb 5.409 | N/A | N/A | N/A | N/A
003303 | Diskmaskin | 003303 | GW0160 | 2.8604e+19 | N/A | N/A | Positivt labb 5.575 | N/A | N/A | N/A | N/A
003045 | Diskmaskin | 003045 | GW0160 | 2.8604e+19 | N/A | N/A | Negativt labb 5.409 | N/A | N/A | N/A | N/A
Infinity-II | Genexpert Infinity-80 | 821122 | Genexpert Infinity-80 | N/A | In-Process | UM0020 Cepheid AB | FI labb 5.544 | Mar-23 | 2 gång/år | N/A | N/A | N/A | N/A
Infinity-VII | Genexpert Infinity-80 | 807363 | Genexpert Infinity-80 | N/A | In-Process | UM0020 Cepheid AB | FI labb 5.544 | Aug-23 | 2 gång/år | N/A | N/A | N/A | N/A
Ugn+pump - används ej
000405 | Vakuumugn | 000405 | VD53 | 11-04137 | N/A | N/A | Korridor | N/A | N/A | 2025-06-01 | 1 gg/år
000408 | Vakuumugn | 000408 | VD53 | 13-05094 | N/A | N/A | Positivt lab 5.576 | N/A | N/A | 2025-06-01 | 1 gg/år
000407 | Pump
000409 | Pump
002289 | Vakuumugn |  | VD53
002333 | Pump
Timer Nr. 22 | Timer | 003757 | EU 609-0206 | D2421661 |  |  | Negativt labb 5.409 | Mar-26 | Timers byts ut efter att kalibreringsdatum utgått | N/A | N/A
Timer Nr. 23 | Timer | 003758 | EU 609-0206 | E2423934 |  |  | Positivt lab 5.576 | Mar-26 | Timers byts ut efter att kalibreringsdatum utgått | N/A | N/A

## Funktionella Error
Minor Functional | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error | Description
| 022025VW_08_0_07+ | 1211744843 | Negative Control 1 | 847922 | 660764 | 2025-10-08T06:24:19+0200 | Aborted | ERROR | 23.2 | 2008 | Max Pressure | Error 2008: Syringe pressure reading of 100.1 PSI exceeds the protocol limit of 100.0 PSI
[10 empty] | 2009 | Max Pressure
| 040925TN_10_1_12X | 1021846357 | Positive Control 1 | 807363 | 880248 | 2025-12-11T20:38:46+0100 | Aborted | ERROR | 41.4 | 2125 | Sample Adequacy | Error 2125: Termination Error - Insufficient Volume: 30, 50, 0, 41
[9 empty] | 90 |  | Max P > 90 PSI
RIF Resistance Indeterminate
| 050525Masa_06_1_14X | 1218982711 | Positive Control 1 | 802069 | 865290 | 2025-11-23T21:13:42+0100 | Aborted | ERROR | 37 | 2096 | Assay Specific Termination Errors | Error 2096: Assay-Specific Termination Error #1: 20, 30, 0, 0
| 061325ELIF_01_1_12X | 218239329 | Specimen | 750172 | 210003099 | 2025-11-20T13:55:54+0100 | Aborted | ERROR | 50 | 2097 | Assay Specific Termination Errors | Error 2097: Assay-Specific Termination Error #2: 46, 7, 1, 0
| 082622VW_04_0_05X | 1189642120 | Specimen | 803094 | 797398 | 2016-05-25 13:13 | Aborted | NO RESULT | 28.4 | 2037 | Cartridge Integrity Failure | Error 2037: The cartridge integrity test failed at valve position 1390. The pressure change of -0.6 PSI did not exceed the requirement of 4.0 PSI. The pressure increased from 3.5 PSI to 3.0 PSI during the test
| 082622VW_05_0_04X | 1189644318 | Specimen | 847922 | 791277 | 2016-05-25 14:26 | Aborted | ERROR | 36 | 5007 | Probe Check | Error 5007: [CT1] probe check failed. Probe check value of 17 for reading number 2 was below the minimum of 110¦Error 5007: [SPC] probe check failed. Probe check value of 19 for reading number 2 was below the minimum of 119
| NEG_07_0_08+ | 600571765 | Specimen | 807363 | 790890 | 2025-08-03T01:52:16+0200 | Aborted | ERROR | 18.7 | 5017 | Probe Check | Error 5017: [SPC] probe check failed. Probe check value of 0 for reading number 2 was below the valid level of 5¦Error 5017: [SAC] probe check failed. Probe check value of 0 for reading number 2 was below the valid level of 5¦Error 5007: [GBS] probe check failed. Probe check value of 5 for reading number 2 was below the minimum of 70
| 022025VW_00_0_08+ | 1211742925 | Negative Control 1 | 847922 | 775845 | 2025-10-08T04:24:46+0200 | Aborted | ERROR | 63.9 | 5006 | Probe Check | Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 1 was above the maximum of 223¦Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 3 was above the maximum of 223
[10 empty] | 5008 | Probe Check
[10 empty] | 5009 | Probe Check
[10 empty] | 5018 | Probe Check
[10 empty] | 5019 | Probe Check
SPC Ct=0
| 011922DL_07_0_07+ | 218569058 | Negative Control 1 | 847922 | 897302 | 2025-11-22T01:57:44+0100 | Done | INVALID | 50 |  | Invalid
[10 empty] | 5001 | Curve Fit Error
[10 empty] | 5002 | Curve Fit Error
[10 empty] | 5003 | Curve Fit Error
[10 empty] | 5004 | Curve Fit Error
[10 empty] | 5005 | Curve Fit Error
| 051625MMM_02_1_11X | 618195071 | Positive Control 1 | 807363 | 781149 | 2025-11-18T08:43:26+0100 | Aborted | ERROR | 32.5 | 5011 | Signal Loss Detection | Error 5011: Signal loss detected in the amplification curve for analyte [mec]. 44.6 decrease in signal with 22.2% decrease at cycle 23.
| 042325MMM_04_0_05X | 922179743 | Negative Control 1 | 847922 | 635023 | 2025-12-12T20:42:58+0100 | Done | INVALID | 55.1 | 5015 | Curve Fit Error | Error 5015: Failed to verify valid background slope for analyte [IQS-H]. The absolute value of the slope of 2.2 was above the maximum of 2.0
[10 empty] | 5016 | Curve Fit Error
[10 empty] | Any Error | Delamination
Major Functional | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error | Description
Instrument fel | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error | Description
| C15NOV23D_03_1_14X | 614338042 | Positive Control 1 | 839032 | 656159 | 2025-10-26T04:57:00+0100 | Aborted | ERROR | 5.8 | 2005 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI
| C15NOV23D_03_1_15+ | 614338040 | Positive Control 1 | 839032 | 210059038 | 2025-10-26T04:57:26+0100 | Done | ERROR | 25 | 1002 | Error 1002: The temperature difference of 5.3 ºC exceeds the limit of 5.0 ºC. The temperatures for heaters A and B are 97.3 ºC and 92.0 ºC
| C15NOV23D_03_1_16+ | 614338465 | Positive Control 1 | 839032 | 210012379 | 2025-10-26T04:57:53+0100 | Done | ERROR | 60.3 | 1001 | Error 1001: The actual temperature 61.6 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.5 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.4 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 61.7 ºC has drifted too far away from the setpoint of 61.0 ºC
Ersättningar | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error | Description | Misquantitation
| 020625VW_09_1_14A+ | 12751975 | Specimen | 807363 | 804538 | 2025-08-04T00:56:44+0200 | Done | Toxigenic C.diff POSITIVE;027 PRESUMPTIVE POS | 25 |  |  | False Result (False Negative or False Positive)
| NEG_05_0_02AX | 907634439 | Specimen | 847922 | 665815 | 2025-09-18T17:28:24+0200 | Done | MTB NOT DETECTED | 11
| 061325ELIF_03_1_11AX | 218240307 | Specimen | 802069 | 914262 | 2025-11-20T16:59:56+0100 | Done | Toxigenic C.diff POSITIVE;027 PRESUMPTIVE POS | 14
Instrument fel | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error | Description

## Infinity GX
Utrustningsfilen
NEG | [7 empty] | Pos
INSTRUMENT | [7 empty] | INSTRUMENT
GeneXpert / Infinity | GX5 | GX6 | Infinity-VI | Infinity-VIII |  |  |  | GeneXpert / Infinity | GX1 | GX2 | GX3 | GX7 | Infinity-I | Infinity-III | Infinity-V
Calibration Due Date | 2026-04-25 | 2026-04-25 | 2026-05-25 | 2025-01-26 |  |  |  | Calibration Due Date | Mar-26 | Mar-26 | Mar-26 | Mar-26 | Jun-26 | Apr-26 | 2026-01-01
PIPETTER | [7 empty] | PIPETTER
Pipette Information | Nr. 6 | Nr. 7 | Nr. 23 | Nr. 32 | Nr. 33 | Nr. 35 | Nr. 36 | Pipette Information | Nr. 3 | Nr. 4 | Nr. 13 | Nr. 17 | Nr. 18 | Nr. 19 | Nr. 20
Calibration Due Date | 2026-02-25 | 2026-02-25 | 2026-03-25 | 2026-01-01 | 2026-01-01 | 2026-02-25 | 2026-02-25 | Calibration Due Date | 2026-01-01 | 2025-12-25 | 2026-03-25 | 2026-01-25 | 2026-03-25 | 2026-03-25 | 2026-03-25
[8 empty] | Pipette Information | Nr. 21 | Nr. 22 | Nr. 24 | Nr. 25 | Nr. 26 | Nr. 34 | Nr. 37
[8 empty] | Calibration Due Date | 2026-03-25 | 2026-01-25 | 2026-03-25 | 2026-03-25 | 2026-03-25 | 2026-02-01 | Sep-25
Uppdaterad: 2025 12 08 / JESP
PIPETTER hämtade från 2099 09 09 GXMTB RIF ULTRA 92303 Worksheet D25862.xlsx
Cepheid ID | Nr. 35 |  | Nr. 21 |  | Nr. 22 |  | N/A | N/A |  | N/A
Calibration Due Date | 2025-11-25 |  | 2025-12-25 |  | 2026-01-25 |  | N/A | N/A |  | N/A
INSTRUMENT hämtade från 2099 09 09 GXMTB RIF ULTRA 92303 Worksheet D25862.xlsx
Cepheid ID | Infinity-I |  | Infinity-V |  | Infinity-VIII |  | Infinity-VI | N/A |  | N/A
Calibration Due Date | 2025-12-25 |  | 2026-01-01 |  | 2026-01-01 |  | 2025-11-25 | N/A |  | N/A
Cepheid ID | N/A |  | N/A |  | N/A |  | N/A | N/A |  | N/A
Calibration Due Date | N/A |  | N/A |  | N/A |  | N/A | N/A |  | N/A

## Information2
Information2 – QC Rule Summary
CSV | 2099 09 09 GXMTB RIF ULTRA 92303 Tests Summary.csv | Generated | 2025-12-28 18:12 | Version: v45.1.0
Assay (canonical) | Xpert MTB-RIF Ultra | Assay Version | 4 | Lot | 92303, 92302 | WorkCenter |  | Match source | Profiles
Row count | 209 | Unique SampleID | 208 | Unique CartridgeSN | 207
KPI / Summary
Findings: ERROR / WARN / INFO | 1 / 2 / 1
Error rows (affected tests) | 2
Distinct error codes | 0
Baseline (expected/actual/delta) | 210 | 209 | -1
Max Pressure (max) | 33.9
Max Pressure >= Warn / Fail | 0 | 0
Status counts (Done/Error/Invalid/Aborted/Incomplete) | 208/0/0/1/0
Findings
Severity | RuleId | Title | Message | Count | Example | Evidence
Error | IDENTITY_REAGENT_LOT_ID | ReagentLotId mismatch (expected '92303') | ReagentLotId mismatch (expected '92303') | 2 | C15NOV23D_07_1_16+
Warn | DUPLICATE_SAMPLEID | Duplicate SampleId detected | Duplicate SampleId detected | 0 |  | Count=1
Warn | DUPLICATE_CARTRIDGESN | Duplicate CartridgeSN detected | Duplicate CartridgeSN detected | 0 |  | Count=2
Info | COLUMN_MISSING_WORKCENTER | Column missing: WorkCenter | Column missing: WorkCenter | 0
Error Summary
Code | Name | Classification | Count | Example SampleID
2037 | 2037 | FUNCTIONAL | 1 | NEG_00_0_04X
5015 | 5015 | FUNCTIONAL | 1 | NEG_08_0_07+
Affected Tests (Warn/Error)
Severity | RuleId | Sample ID | Cartridge S/N | Module S/N | Test Type | Status | Test Result | Max Pressure (PSI) | Error | ErrorCode | WorkCenter | Row#
Error | IDENTITY_REAGENT_LOT_ID | C15NOV23D_07_1_16+ | 614350191 | 638845 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 30.1 |  |  |  | 82
Error | IDENTITY_REAGENT_LOT_ID | NEG_10_0_10+ | 614390273 | 652001 | Negative Control 1 | Done | MTB NOT DETECTED | 18.2 |  |  |  | 218

## MTB ULTRA
MTB ULTRA
P/N | Artikel | Lotnr. | Utgångsdatum | Förvaringsplats | Antal i lager | Senast inventering | SIGN
001-1923 | INTROL TB1081MDR2 | B27FEB25A | 2027-02-28 | IPT kylskåp | 5 Lådor | 2025-11-28 | JESP
001-1923 |  | N/A | N/A |  | N/A | N/A | N/A
001-1923 |  | N/A | N/A |  | N/A | N/A | N/A
001-1923 |  | N/A | N/A |  | N/A | N/A | N/A
P/N | Artikel | Lotnr. | Utgångsdatum | Förvaringsplats | Antal i lager | Senast inventering | SIGN
001-1968 | INTROL TBWT-04 MTB Ultra | B01NOV24B | 2026-11-30 | IPT kylskåp | 5.5 lådor | 2025-11-24 | ELS
001-1968 |  | N/A | N/A |  | N/A | N/A | N/A
001-1968 |  | N/A | N/A |  | N/A | N/A | N/A
001-1968 |  | N/A | N/A |  | N/A | N/A | N/A
MTB - TET ¦ 500-0995
P/N 500-0995 MTB TET Buffer (exp 6 mån) 500-0995, BUFFER, TET, 50_0.1_0.1
Lotnummer |  |  | Utgångsdatum |  | Volym, L | Senast inventering | SIGN
1001510545 |  |  | 2026-05-14 |  | 25L | 2025-11-21 | ELS
1001503562 |  |  | 2026-03-30 |  | 10L | 2025-11-24 | vdao
N/A |  |  | N/A |  | N/A | N/A | N/A
N/A |  |  | N/A |  | N/A | N/A | N/A
MTB - SR ¦ 500-2608
P/N 500-2608 MTB SR (exp 4 mån) 500-2608, IPT BULK REAGENT, SAMPLE TREATMENT, MTB
Lotnummer |  |  | Utgångsdatum |  | Volym, L | Senast inventering | SIGN
1001508989 |  |  | 2026-03-06 |  | 50L | 2025-11-21 | ELS
1001505209 |  |  | 2026-02-10 |  | 30 | 2025-11-18 | JESP
N/A |  |  | N/A |  | N/A | N/A | N/A
N/A |  |  | N/A |  | N/A | N/A | N/A

## SharePoint Info
SharePoint Information
Work Center | ROBAL 06
Order# | 1110580
SAP Batch# | 1001506046
LSP | 92303
Material | GXMTB/RIF-ULTRA-50
BBD/SLED | 2027-04-11
ROBAL - Actual start date/time | 2025-10-25 22:40
Sample Reagent use | 1001501891, 1001502674
Sample Reagent P/N | 700-8870
Order quantity | 29000
ROBAL - Till Packning | 27958
IPT - Testing Finalized | 2025-10-26 23:00
MES Order | JA

## ResultRules
Tab | TestTypePattern | Idx | ExpectedResultRegex | Comment
HIV VL | ^Negative Control 1$ | 0 | ^HIV-1 NOT DETECTED$ | HIV VL NC1 – odetekterbar, allt annat (utan tekniskt error) = FN
HIV VL | ^Positive Control 1$ | 1 | ^HIV-1 DETECTED \d+ copies/mL \(log 2\.\d+\)$ | HIV VL PC1 – låg nivå runt 10^2 kopior (log 2.x)
HIV VL | ^Positive Control 2$ | 2 | ^HIV-1 DETECTED [0-9.]+E0?5 copies/mL \(log 5\.\d+\)$ | HIV VL PC2 – hög nivå runt 10^5 kopior (log 5.x)
HIV VL XC | ^Negative Control 1$ | 0 | ^HIV-1 NOT DETECTED$ | HIV VL XC NC1 – odetekterbar
HIV VL XC | ^Positive Control 1$ | 1 | ^HIV-1 DETECTED \d+ copies/mL \(log 2\.\d+\)$ | HIV VL XC PC1 – låg viremi (log 2.x)
HIV VL XC | ^Positive Control 2$ | 2 | ^HIV-1 DETECTED [0-9.]+E0?5 copies/mL \(log 5\.\d+\)$ | HIV VL XC PC2 – hög viremi (log 5.x)
HPV | ^Negative Control 1$ | 0 | ^INVALID$ | HPV NC1 – designat INVALID (v1 & v2)
HPV | ^Positive Control [12]$ | 1 | ^(?:HR HPV POS¦HPV 16 POS;HPV 18_45 POS;OTHER HR HPV POS)$ | HPV PC1/PC2 – positiva kontroller (idx1)
HPV | ^Positive Control [12]$ | 2 | ^(?:HR HPV POS¦HPV 16 POS;HPV 18_45 POS;OTHER HR HPV POS)$ | HPV PC1/PC2 – positiva kontroller (idx2)
MRSA NxG | ^Negative Control 1$ | 0 | ^MRSA NOT DETECTED$ | MRSA NxG NC1 – alltid negativ
MRSA NxG | ^Positive Control [12]$ | 1 | ^MRSA DETECTED$ | MRSA NxG PC1/PC2 – positiva kontroller (idx1/2)
MRSA NxG | ^Positive Control [12]$ | 2 | ^MRSA DETECTED$ | MRSA NxG PC1/PC2 – positiva kontroller (idx1/2)
MRSA SA | ^Specimen$ | 0 | ^MRSA NEGATIVE;SA NEGATIVE$ | SA Nasal Specimen idx0 – dubbelnegativ panel
MRSA SA | ^Specimen$ | 1 | ^MRSA POSITIVE;SA POSITIVE$ | SA Nasal Specimen idx1 – dubbelpositiv panel
CARBA R | ^Negative Control 1$ | 0 | ^IMP NOT DETECTED;VIM NOT DETECTED;NDM NOT DETECTED;KPC NOT DETECTED;OXA48 NOT DETECTED$ | Carba-R NC1 – alla fem gener NOT DETECTED
CARBA R | ^Positive Control 1$ | 1 | ^IMP DETECTED;VIM DETECTED;NDM DETECTED;KPC DETECTED;OXA48 DETECTED$ | Carba-R PC1 – full house positiv
CARBA R | ^Positive Control 2$ | 2 | ^IMP DETECTED;VIM DETECTED;NDM DETECTED;KPC DETECTED;OXA48 DETECTED$ | Carba-R PC2 – full house positiv
FLUVID ¦ FLUVID+ | ^Negative Control 1$ | 0 | ^SARS-CoV-2 NEGATIVE;Flu A NEGATIVE;Flu B NEGATIVE;RSV NEGATIVE$ | XP3/FLUVID NC1 – alla fyra targets NEGATIVE
FLUVID ¦ FLUVID+ | ^Positive Control 1$ | 1 | ^SARS-CoV-2 POSITIVE;Flu A POSITIVE;Flu B POSITIVE;RSV POSITIVE$ | XP3/FLUVID PC1 – alla fyra targets POSITIVE
CTNG ¦ CTNG JP | ^Specimen$ | 0 | ^(?:CT NOT DETECTED;NG NOT DETECTED¦CT NOT DETECTED)$ | CT/NG Specimen idx0 – dubbelnegativ panel
CTNG ¦ CTNG JP | ^Specimen$ | 1 | ^(?:CT DETECTED;NG DETECTED¦CT DETECTED)$ | CT/NG Specimen idx1 – dubbelpositiv panel
MTB ULTRA | ^Negative Control 1$ | 0 | ^MTB NOT DETECTED$ | MTB Ultra NC1 – ingen MTB (Trace = FP-kandidat)
MTB ULTRA | ^Positive Control 1$ | 1 | ^MTB DETECTED (?:LOW¦VERY LOW¦MEDIUM¦HIGH);RIF Resistance (?:DETECTED¦INDETERMINATE)$ | MTB Ultra PC1 – MTB DETECTED, RIF DETECTED/INDETERMINATE
MTB ULTRA | ^Positive Control 2$ | 2 | ^MTB DETECTED (?:LOW¦VERY LOW¦MEDIUM¦HIGH);RIF Resistance NOT DETECTED$ | MTB Ultra PC2 – MTB DETECTED, RIF NOT DETECTED
HIV QA | ^Specimen$ | 0 | ^HIV-1 NOT DETECTED$ | HIV Qual Specimen idx0 – negativ panel
HIV QA | ^Specimen$ | 1 | ^HIV-1 DETECTED$ | HIV Qual Specimen idx1 – positiv panel
C.DIFF ¦ C.DIFF JP | ^Specimen$ | 0 | ^Toxigenic C\.diff NEGATIVE;027 PRESUMPTIVE NEG$ | C.diff Specimen idx0 – NEG; panel-negativ
C.DIFF ¦ C.DIFF JP | ^Specimen$ | 2 | ^Toxigenic C\.diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC2/PC3 – POS; panel-positiv
C.DIFF ¦ C.DIFF JP | ^Specimen$ | 1 | ^Toxigenic C\.diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC1 – majoritet POS; enstaka NEG ses som FN/edge-cases
Tab | TestTypePattern | Idx | ExpectedResultRegex | Comment
C.DIFF | ^Specimen$ | 0 | ^Toxigenic C\.?diff NEGATIVE;027 PRESUMPTIVE NEG$ | C.diff Specimen – negativ panel
C.DIFF | ^Positive Control 1$ | 1 | ^Toxigenic C\.?diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC1 – majoritet POS
C.DIFF | ^Positive Control [23]$ | 2 | ^Toxigenic C\.?diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC2/PC3 – POS
C.DIFF CN | ^Specimen$ | 0 | ^Toxigenic C\.?diff NEGATIVE;027 PRESUMPTIVE NEG$ | C.diff Specimen – negativ panel
C.DIFF CN | ^Positive Control 1$ | 1 | ^Toxigenic C\.?diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC1 – majoritet POS
C.DIFF CN | ^Positive Control [23]$ | 2 | ^Toxigenic C\.?diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC2/PC3 – POS
C.DIFF JP | ^Specimen$ | 0 | ^Toxigenic C\.?diff NEGATIVE;027 PRESUMPTIVE NEG$ | C.diff Specimen – negativ panel
C.DIFF JP | ^Positive Control 1$ | 1 | ^Toxigenic C\.?diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC1 – majoritet POS
C.DIFF JP | ^Positive Control [23]$ | 2 | ^Toxigenic C\.?diff POSITIVE;027 PRESUMPTIVE POS$ | C.diff PC2/PC3 – POS
CARBA-R |  |  |  | Öppet för mig att fylla i
CTNG | ^Specimen$ | 0 | ^(?:CT NOT DETECTED;NG NOT DETECTED¦CT NOT DETECTED)$ | CT/NG Specimen – dubbelnegativ panel
CTNG | ^Specimen$ | 1 | ^(?:CT DETECTED;NG DETECTED¦CT DETECTED)$ | CT/NG Specimen – dubbelpositiv panel
EBOLA |  |  |  | Öppet för mig att fylla i
FLU NxG | ^Negative Control 1$ | 0 | ^SARS-CoV-2 NEGATIVE;Flu A NEGATIVE;Flu B NEGATIVE;RSV NEGATIVE$ | FLUVID NC1 – alla targets negativa
FLU NxG | ^Positive Control 1$ | 1 | ^SARS-CoV-2 POSITIVE;Flu A POSITIVE;Flu B POSITIVE;RSV POSITIVE$ | FLUVID PC1 – alla targets positiva
FLU RSV | ^Negative Control 1$ | 0 | ^SARS-CoV-2 NEGATIVE;Flu A NEGATIVE;Flu B NEGATIVE;RSV NEGATIVE$ | FLUVID NC1 – alla targets negativa
FLU RSV | ^Positive Control 1$ | 1 | ^SARS-CoV-2 POSITIVE;Flu A POSITIVE;Flu B POSITIVE;RSV POSITIVE$ | FLUVID PC1 – alla targets positiva
FLUVID |  |  |  | Öppet för mig att fylla i
FLUVID+ |  |  |  | Öppet för mig att fylla i
GBS |  |  |  | Öppet för mig att fylla i
HBV VL |  |  |  | Öppet för mig att fylla i
HCV VL |  |  |  | Öppet för mig att fylla i
HIV QA | ^Specimen$ | 0 | ^HIV-1 NOT DETECTED$ | HIV Qual Specimen – negativ
HIV QA | ^Specimen$ | 1 | ^HIV-1 DETECTED$ | HIV Qual Specimen – positiv
HIV QA XC |  |  |  | Öppet för mig att fylla i
HIV VL | ^Negative Control 1$ | 0 | ^HIV-1 NOT DETECTED$ | HIV VL NC1 – odetekterbar
HIV VL | ^Positive Control 1$ | 1 | ^HIV-1 DETECTED \d+ copies/mL \(log 2\.\d+\)$ | HIV VL PC1 – låg viremi
HIV VL | ^Positive Control 2$ | 2 | ^HIV-1 DETECTED [0-9.]+E0?5 copies/mL \(log 5\.\d+\)$ | HIV VL PC2 – hög viremi
HIV VL XC | ^Negative Control 1$ | 0 | ^HIV-1 NOT DETECTED$ | HIV VL NC1 – odetekterbar
HIV VL XC | ^Positive Control 1$ | 1 | ^HIV-1 DETECTED \d+ copies/mL \(log 2\.\d+\)$ | HIV VL PC1 – låg viremi
HIV VL XC | ^Positive Control 2$ | 2 | ^HIV-1 DETECTED [0-9.]+E0?5 copies/mL \(log 5\.\d+\)$ | HIV VL PC2 – hög viremi
HPV | ^Negative Control 1$ | 0 | ^INVALID$ | HPV NC1 – INVALID
HPV | ^Positive Control [12]$ | 1 | ^(?:HR HPV POS¦HPV 16 POS;HPV 18_45 POS;OTHER HR HPV POS)$ | HPV PC1/PC2 – POS
HVC VL FS |  |  |  | Öppet för mig att fylla i
MRSA NxG | ^Negative Control 1$ | 0 | ^MRSA NOT DETECTED$ | MRSA NxG NC1 – negativ
MRSA NxG | ^Positive Control [12]$ | 1 | ^MRSA DETECTED$ | MRSA NxG PC1/PC2 – positiv
MRSA NxG | ^Positive Control [12]$ | 2 | ^MRSA DETECTED$ | MRSA NxG PC1/PC2 – positiv
MRSA SA | ^Specimen$ | 0 | ^MRSA NEGATIVE;SA NEGATIVE$ | MRSA SA Specimen – dubbelnegativ
MRSA SA | ^Specimen$ | 1 | ^MRSA POSITIVE;SA POSITIVE$ | MRSA SA Specimen – dubbelpositiv
MRSA SA CN | ^Specimen$ | 0 | ^MRSA NEGATIVE;SA NEGATIVE$ | MRSA SA Specimen – dubbelnegativ
MRSA SA CN | ^Specimen$ | 1 | ^MRSA POSITIVE;SA POSITIVE$ | MRSA SA Specimen – dubbelpositiv
MRSA SA JP | ^Specimen$ | 0 | ^MRSA NEGATIVE;SA NEGATIVE$ | MRSA SA Specimen – dubbelnegativ
MRSA SA JP | ^Specimen$ | 1 | ^MRSA POSITIVE;SA POSITIVE$ | MRSA SA Specimen – dubbelpositiv
MTB CN |  |  |  | Öppet för mig att fylla i
MTB JP |  |  |  | Öppet för mig att fylla i
MTB RIF |  |  |  | Öppet för mig att fylla i
MTB ULTRA | ^Negative Control 1$ | 0 | ^MTB NOT DETECTED$ | MTB Ultra NC1 – ingen MTB
MTB ULTRA | ^Positive Control 1$ | 1 | ^MTB DETECTED (?:LOW¦VERY LOW¦MEDIUM¦HIGH);RIF Resistance (?:DETECTED¦INDETERMINATE)$ | MTB Ultra PC1 – MTB detekterad, RIF Det./Indet.
MTB ULTRA | ^Positive Control 2$ | 2 | ^MTB DETECTED (?:LOW¦VERY LOW¦MEDIUM¦HIGH);RIF Resistance NOT DETECTED$ | MTB Ultra PC2 – MTB detekterad, RIF ND
MTB XDR |  |  |  | Öppet för mig att fylla i
NORO |  |  |  | Öppet för mig att fylla i
SARS-COV-2 |  |  |  | Öppet för mig att fylla i
SARS-COV-2 Plus |  |  |  | Öppet för mig att fylla i
SARS-COV-3 |  |  |  | Öppet för mig att fylla i
SARS-COV-4 |  |  |  | Öppet för mig att fylla i
STREP A |  |  |  | Öppet för mig att fylla i
VAN A/(B) |  |  |  | Öppet för mig att fylla i
[5 empty] | Status | Test Result | Error | Description
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 2008 | Max Pressure
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 2009 | Max Pressure
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 2125 | Sample Adequacy
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID |  | Max P > 90 PSI
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 2096 | Assay Specific Termination Errors
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 2097 | Assay Specific Termination Errors
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 2037 | Cartridge Integrity Failure
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5007 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5017 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5006 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5008 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5009 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5018 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5019 | Probe Check
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID |  | SPC Ct=0
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID |  | Invalid
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5001 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5002 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5003 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5004 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5005 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5011 | Signal Loss Detection
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5015 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | 5016 | Curve Fit Error
[5 empty] | Aborted / Done | ERROR / NO RESULT / INVALID | Any Error | Delamination
[5 empty] | HPV NC1 Invalid | For HPV assays, Negative Control 1 is expected to be INVALID by design and is not considered a functional failure.
[5 empty] | Max Pressure limit | Max Pressure (PSI) must never exceed 90. Any test with Max Pressure > 90 PSI is classified as functional pressure failure (no re-test).
[5 empty] | Invalid w/o error (non-HPV) | Invalid result without an error code combined with SPC Ct=0 and non-HPV assay (and where expected category is not INVALID) is classified as Minor Functional (true functional failure).
[5 empty] | Instrument error vs Minor Functional | Error codes listed explicitly in ErrorBank are treated as Minor Functional (no re-test). All other numeric error codes are treated as Instrument Error and should result in cartridge replacement / re-test.
[5 empty] | False Result (False Positive or false Negative) | When an specific expected target deviate from Assay rules

## ErrorCode
ErrorCode | Name | Group | Classification | Description | GeneratesRetest
5001 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5002 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5003 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5004 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5005 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5015 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5016 | Curve Fit Error | CurveFit | FUNCTIONAL_MINOR_CURVEFIT | Curve fit error – true functional failure, no re-test | No
5007 | Probe Check Low | ProbeCheckLow | FUNCTIONAL_MINOR_PROBECHECK_LOW | Probe Check Low – true functional failure, no re-test | No
5017 | Probe Check Low | ProbeCheckLow | FUNCTIONAL_MINOR_PROBECHECK_LOW | Probe Check Low – true functional failure, no re-test | No
5019 | Probe Check Low | ProbeCheckLow | FUNCTIONAL_MINOR_PROBECHECK_LOW | Probe Check Low – true functional failure, no re-test | No
5006 | Probe Check High | ProbeCheckHigh | FUNCTIONAL_MINOR_PROBECHECK_HIGH | Probe Check High – true functional failure, no re-test | No
5018 | Probe Check High | ProbeCheckHigh | FUNCTIONAL_MINOR_PROBECHECK_HIGH | Probe Check High – true functional failure, no re-test | No
2037 | CIT | CIT/SLD | FUNCTIONAL_MINOR_CIT_SLD | CIT – true functional failure, no re-test | No
5011 | SLD | CIT/SLD | FUNCTIONAL_MINOR_CIT_SLD | SLD – true functional failure, no re-test | No
2008 | Pressure Abort | Pressure | FUNCTIONAL_MINOR_PRESSURE | Pressure abort; protocol limit exceeded (also if Max Pressure > 90 PSI) | No
2096 | Sample Volume Adequacy | SVA | FUNCTIONAL_MINOR_SVA | Sample Volume Adequacy failure – true functional failure, no re-test | No
2097 | Sample Volume Adequacy | SVA | FUNCTIONAL_MINOR_SVA | Sample Volume Adequacy failure – true functional failure, no re-test | No
2125 | Sample Volume Adequacy | SVA | FUNCTIONAL_MINOR_SVA | Sample Volume Adequacy failure – true functional failure, no re-test | No
| Max Pressure > 90 PSI | Pressure | FUNCTIONAL_MINOR_PRESSURE | Max Pressure (PSI) strictly > 90 – functional failure, no re-test | No
| Invalid w/o error (non-HPV, SPC Ct=0) | InvalidNoCode | FUNCTIONAL_MINOR_INVALID | Invalid result without error code, SPC Ct=0, non-HPV assay – true functional failure, no re-test | No
| Other error codes | Other | INSTRUMENT_ERROR | All other error codes – classified as instrument error, cartridge re-test/replacement required | Yes
| Delamination + error code | Delamination | FUNCTIONAL_MINOR_DELAM | Any error code combined with delamination suffix (D01, D2A, ... D11) – true functional failure, no re-test | No
| RIF Resistance Indeterminate | MTB | FUNCTIONAL_MINOR_MTB_ASSAYS_POS | If any MTB Assay expected to be positive has "Indeterminate" in Test Result then it's a true minor functional | No

## ExpectedResults
Assay
Xpert C.difficile G3
Xpert C.difficile G3
Xpert C.difficile G3
Xpert C.difficile G3

## GeneralRules
Rule | Description
Idx mapping (Controls) | idx 0 = Negative Control 1; idx 1 = Positive Control 1; idx 2 = Positive Control 2; idx 3 = Positive Control 3; idx 4 = Positive Control 4; idx 5 = Positive Control 5. Assays with Test Type = 'Specimen' are exceptions where idx represents panel layout rather than control index.
Sample ID structure | General pattern: PREFIX_BAG_IDX_POS+SUFFIX, where BAG = 00–10, IDX = 0–5, POS = 01–20. PREFIX = NEG/POS or cell name; SUFFIX can include A/AA/AAA for replacements and D01..D11 for delamination on specific chimneys.
HPV NC1 Invalid | For HPV assays (Tab=HPV), Negative Control 1 (idx 0) is expected to be INVALID by design and is not considered a functional failure.
Max Pressure limit | Max Pressure (PSI) must never exceed 90. Any test with Max Pressure > 90 PSI is classified as functional pressure failure (no re-test).
Invalid w/o error (non-HPV) | Invalid result without an error code combined with SPC Ct=0 and non-HPV assay (and where expected category is not INVALID) is classified as Minor Functional (true functional failure).
Instrument error vs Minor Functional | Error codes listed explicitly in ErrorBank are treated as Minor Functional (no re-test). All other numeric error codes are treated as Instrument Error and should result in cartridge replacement / re-test.

## Blad1
AssaySlang | Part/Cat. No. | Assay | DocNumber | SampleType | SampleTypeSize | Bag00Role | Bag00RepRange | Bag01-10RepRangeCode0 | ExpectedSampleTypeCode0 | ExpectedSampleTypeCode1 | OtherBagsRepRange | BagsUsed | ExpectedSampleTypeCode0 | ExpectedSampleTypeCode1 | ExpectedSampleTypeCode2 | ExpectedSampleTypeCode4 | ExpectedSampleTypeCode5 | ExpectedTotalPerBatch | ValidSuffixes
MTB RIF |  | Xpert MTB-RIF Assay G4 | D31503 | Specimen |  | _00 | 10 | 01,02,03,04,05,06,07,08,09,10 | 110 |  | 01–10 | 00–10 | 110 | 100 |  |  |  | 210 | X,+,A,AA,AAA,D##
MTB JP |  | Xpert MTB-RIF JP IVD | D66612 | Specimen |  | _00 | 10 |  | 110 |  | 11–20 | 00–10 | 110 |  |  |  |  | 210 | X,+,A,AA,AAA,D##
MTB ULTRA |  | Xpert MTB-RIF Ultra | D25862 |  |  | _00 | 10
MTB XDR |  | Xpert MTB-XDR | D37339 | NC1 |  | _00 | 10 |  | 150 |  | 1–15 | 00–10 | 150 |  |  |  |  | 210 | X,+,A,AA,AAA,D##
FLU RSV |  | Xpress Flu IPT_EAT off | D26120 | PC1 |  | _00 | 10 |  | 30 |  | 16–18 | 01–10 | 30 |  |  |  |  | 210 | X,+,A,AA,AAA,D##
FLU RSV |  | Xpert Xpress Flu-RSV | D26120 | PC2 |  | _00 | 10 |  | 20 |  | 19–20 | 01–10 | 20 |  |  |  |  | 210 | X,+,A,AA,AAA,D##
FLUVID ¦ FLUVID+ |  | Xpert Xpress_SARS-CoV-2_Flu_RSV | D41929 |  |  | _00 | 10
FLUVID ¦ FLUVID+ |  | Xpress SARS-CoV-2_Flu_RSV plus | D47377 |  |  | _00 | 10
SARS-COV-2 Plus |  | Xpert Xpress CoV-2 plus | D48538 |  |  | _00 | 10
CTNG ¦ CTNG JP |  | Xpert CT_CE | D16904 |  |  | _00 | 10
CTNG ¦ CTNG JP |  | Xpert CT_NG | D16904 |  |  | _00 | 10
CTNG ¦ CTNG JP |  | Xpert CT_NG | D66615 |  |  | _00 | 10
C.DIFF ¦ C.DIFF JP | 700-5178 | Xpert C.difficile BT | D37468 |  |  | _00 | 10
C.DIFF ¦ C.DIFF JP | 700-5102 | Xpert C.difficile G3 | D37468 |  |  | _00 | 10
C.DIFF ¦ C.DIFF JP | 700-5102 | Xpert C.difficile G3 | D66613 |  |  | _00 | 10
EBOLA |  | Xpert Ebola EUA | D21938 |  |  | _01 | 10
EBOLA |  | Xpert Ebola CE-IVD | D21938 |  |  | _01 | 10
HPV |  | Xpert HPV v2 HR | D16546 |  |  | _01 | 10
HPV |  | Xpert HPV HR | D16546 |  |  | _01 | 10
HBV VL |  | Xpert HBV Viral Load | D51114 |  |  | _01 | 10
HCV |  | Xpert HCV PQC | D70915 |  |  | _01 | 10
HCV VL |  | Xpert_HCV Viral Load | D51112 |  |  | _01 | 10
HCV VL FS |  | Xpert HCV VL Fingerstick | D51113 |  |  | _01 | 10
HIV VL |  | Xpert_HIV-1 Viral Load | D51110 |  |  | _01 | 10
HIV VL XC |  | Xpert HIV-1 Viral Load XC | D51111 |  |  | _01 | 10
HIV QA |  | Xpert_HIV-1 Qual | D61353 |  |  | _01 | 10
HIV QA XC |  | Xpert HIV-1 Viral Load XC | D79377 |  |  | _01 | 10
SARS-COV-2 |  | Xpert Xpress SARS-CoV-2 CE-IVD | D39525 |  |  | _00 | 10
MRSA SA |  | GXSACOMP-CE-10 | D36872 |  |  | _00 | 10
MRSA SA |  | GXMRSA/SA-SSTI-CE | D36872 |  |  | _00 | 10
MRSA SA |  | GXSACOMP-CN-10 | D36872 |  |  | _00 | 10
MRSA JP |  | Xpert SA Nasal Complete G3 | D66614 |  |  | _00 | 10
MRSA NxG |  | Xpert MRSA NxG | D23916 |  |  | _00 | 10
NORO |  | Xpert Norovirus | D17716 |  |  | _01 | 10
VAN AB |  | Xpert vanA vanB | D55782 |  |  | _01 | 10
GBS |  | Xpert Xpress GBS | D37988 |  |  | _01 | 10
GBS |  | Xpert GBS LB XC | D68621 |  |  | _01 | 10
GBS |  | Xpert Xpress GBS US-IVD | D68620 |  |  | _01 | 10
STREP A |  | Xpert Xpress Strep A | D27089 |  |  | _00 | 10
STREP A |  | Xpert Xpress_Strep A | D27089 |  |  | _00 | 10
CARBA R |  | Xpert Carba-R | D18272 |  |  | _00 | 10
CARBA R |  | Xpert_Carba-R | D18272 |  |  | _00 | 10

## Instrument Error
Failure Code
1001-Temp. Drifted from Setpoint
1002-Heater Temp. Differential Error
1004-Internal Temp. out of range
1005-Optical Signal Too High
1006-Detector Dark Signal Error
1007-Power Supply Voltage Error
1011-Optical Read Time Too Long
1017-Optical Block Temperature out of range
1018-Valve Position Error at End of Test (>/=6 counts in 4.0)
1018-Valve Position Error at End of Test (>/=6 counts in 4.0)
1032 - Module [site name] cannot start test due to internal temperature at f.f degree C above threshold of f.f.
1035 - Module [site name]: Requires Cleaning (nn tests since Plunger Rod was last cleaned). Note: This error message is applicable only for Xpertise 6.2 and above. The message is displayed if more than 1000 tests are created after the last clean.
1091 - Cannot perform cartridge preparation because a run is already active
1092 - Cannot perform cartridge preparation because cartridge preparation is already underway.
1096 - Proceeded To Next Step #1: n, m, p, q (n, m, p, q values are assay specific)
1097 - Proceeded To Next Step #2: n, m, p, q (n, m, p, q values are assay specific)
1098 - Proceeded To Next Step #3: n, m, p, q (n, m, p, q values are assay specific)
1099 - Proceeded To Next Step #4: n, m, p, q (n, m, p, q values are assay specific)
1100 - Proceeded To Next Step #5: n, m, p, q (n, m, p, q values are assay specific)
1125 - Insufficient Volume Error
1606 - Syringe error
1650 - Unable to write to archive file
2001-Command n was not recognized
2002-Could not find calibration information for reporter #n
2003-Module Already Running a Test
2005-Motion of Syringe Drive not Detected
2006-Valve Motion not detected
2007 Unable to calculate Optical Signal
2008 Over Pressure Abort
2009-Under Pressure Abort (not typically seen or used in the field)
2011-Unable to Initialize Pressure
2012-Valve Inaccuracy Error
2014-Thermistor Value out of Range
2016-Unable to Find Valve Home
2017-Door Latch Sensor Error
2018-Door Closed at Cartridge Load
2019 - During test playback, an unrecognized datagram was encountered
2021 - Unable to process assay. An unknown primitive #n was encountered
2022-Failure to get to Temp. Setpoint
2023 GX firmware corruption. Unable to continue the test
2024-Ultrasonic Horn Failure
2025-Cartridge Load Error (2 modes)
2026-Ultrasonic Overcurrent
2029 - Cartridge Integrity Test Failure (4.0 and lower)
2037- The Cartridge Integrity Test Failed (4.0b and higher)
2030-Firmware has a Defect
2031-Firmware has a Defect
2032-Ultrasonic Tuning Error
2034-Opt Det Unable to Reach Expected Value
2035-Ultrasonic Failure Occurred
2036 - Module door failed to latch
2054 - Internal Communication Error. Contact technical support.
2096 - Assay-Specific Termination Error #1: n, m, p, q (n, m, p, q values are assay specific)
2097 - Assay-Specific Termination Error #2: n, m, p, q (n, m, p, q values are assay specific)
2098 - Assay-Specific Termination Error #3: n, m, p, q (n, m, p, q values are assay specific)
2099 - Assay-Specific Termination Error #4: n, m, p, q (n, m, p, q values are assay specific)
2100 - Assay-Specific Termination Error #5: n, m, p, q (n, m, p, q values are assay specific)
2120 - Module X lost communication while module was idle
2121 - Module X lost communication before module door was latched
2122 - Module X lost communication while starting test, test aborted
2123 - Module X lost communication while test was running, attempting recovery (these codes are used starting in Dx 4.6)
2124 - Module X communication restored
2125 - Termination Error - Insufficient Volume: n, m, p, q (n, m, p, q values are assay specific)
2126 - Module X was reset
2127 - Module Communication Loss was detected (this code is used before Dx 4.6)
2651 - Protocol command is too large
2652 - Protocol command contains a part with an invalid declared size
2653 - Protocol command data was out of sync
2654 - Protocol command data was not in sequential order
2655 - Protocol command total size was not consistent with the size declared in the header
2656 - Protocol command contains a section larger than was declared
2657 - Protocol command checksum was not correct
2658 - Protocol command header is invalid
2659 - Unsupported protocol command version
2660 - The declared size of the protocol command was not valid
2661 - The declared number of sections in the protocol command was not valid
2662 - Protocol command contains an invalid section type
2663 - Protocol command has duplicated sections
2664 - The provided protocol command sections were not contiguous
2665 - The declared size of a section of the protocol command was not valid
2666 - The declared size of the protocol command does not match the total size of all its section
2667 - Protocol command parameters are not valid
2720 - Failed to download biochip profile: {0}
2721 - Biochip cartridge preparation: {0}"
2722 - Failed to start biochip run: {0}
2723 - Biochip run ended abnormally: {0}, {1}
2724 - Invalid chip thermal calibration parameters were provided
2725 - HPU configuration is invalid.
2727 - Communication with the detector failed with bad message checksum
2729 - Protocol command was not actively running when expected
2731 - Failed to stop protocol command.
2732 - Failed to stop protocol command due to firmware resource issue.
2861 - Attempted to start protocol while a protocol was already active
2862 - MPU configuration is invalid
2863 - Unable to start protocol: sensors failed to initialize.
2864 - MPU health check failed.
2865 - HPU configuration is invalid.
2866 - Unable to start protocol: cartridge chip was not ready
2867 - HPU health check failed
2868 - Unable to start protocol command: cannot coordinate with HPU
2869 - Protocol command was not started because post cartridge preparation steps are being performed.
2951 - Failed to synchronize protocol command instructions.
2962 - Imaging command was out of sync with current state.
2963 - No imaging configuration found for the specified index
2964 - Failed to run setup configuration request
2965 - Setup configuration request received an invalid response
2966 - Failed to run imaging configuration request
2967 - Imaging configuration request. received an invalid response
2968 - IO failure detected with cartridge chip
2972 - Frame interrupt request was late
2973 - Interrupt request status indicates error
2974 - Synchronization error in run setup configuration
2975 - Synchronization error in run setup configuration type 2
2981 - Run aborted due to thermal command scheduling error.
2982 - Run aborted by command from SOM.
2983 - Inter-task commands for control task are out of sync.
2984 - Detected missing or invalid thermal configuration.
2985 - Monitor detected PID has lost control.
2986 - Alert detected and forced abort.
2987 - Imaging failure detected.
2988 - Temperature sensor failed.
2989 - ERS gross violation for temperature
2990 - ERS gross violation for cooling rate.
2992 - Failed to detect cartridge chip after sample preparation; parameters: [{0}] [{1}] [{2}] [{3}].
3000 - Communication buffer overflow error
3200 - Module has insufficient memory
4001-I-CORE Memory Error
4002-GX Memory Error
4003-Ultrasonic Drive Circuitry Failed
4004-Valve Motion Not Detected Error
4006-Syringe Stall Error
4007- Module Syringe Power Rail Error
4008-Power Supply Voltage Error
4009-Heater A Self Test Error
4010-Cooling Self Test Error
4011-Detector Dark Signal Error
4012-Heater B Self Test Error
4013-Valve Move Inaccuracy Error
4014-Optics Self Test Error
4015-Optical Block Temperature Out of Range
4016-Firmware Problems
4017- Thermistor Error
4018-FW Not Compatible with Valve Label
4019 The Optical Ramp Test for LED n resulted in non-monotonic results at DAC setting of nnn
4020 - Module Horn Power Rail Error
4021 - Module servo motion was not detected during cartridge preparation
4022 - Detector LED was not present during cartridge preparation.
4023 - Cartridge preparation timeout error occurred.
4024 - Unable to perform cartridge preparation due to firmware resource issue.
4025 - Cartridge chip ID configuration error detected.
4026 - Invalid cartridge chip trim detected
4027 - Invalid cartridge chip setup detected
4028 - Failed to power cartridge chip
4029 - Failed to detect cartridge chip
4030 - Cartridge chip heater encountered an error
4031 - Failed to connect with cartridge chip
4042 - Module servo motion was not detected during self test.
4043 - Detector LED was not present during self test
4045 - Module health check failed
4142 Module<module name> firmware update failed
4143 Module <module name> firmware version is out of range. Update software
4144 Module <module name> firmware image is not found or firmware image cannot be verified
4145 Module <module name> firmware version is not consistent. Update module firmware
4146 Module <module name> module type is not supported
C- 4808 An ultrasonic horn command failed to be executed. Code {0}
C-4809 The ultrasonic horn failed to find resonance
C-4810 An ultrasonic horn failure occurred with frequency out of range ({0} Hz).
5001 - Unable to verify positive analyte [analyte name] using curve fitting. f.ff value of <parameter name>
5002 - Failed to verify valid amplification curve for analyte [analyte name]. The shape factor of f.ff was below the minimum of f.ff
5003 - Failed to verify valid amplification curve for analyte [analyte name]. The shape factor of f.f was above the maximum of f.f
5004 - Failed to verify valid amplification curve for analyte [analyte name]. The normalized sum of errors of f.f was above the maximum of f.f
5005 - Failed to verify valid amplification curve for analyte [analyte name]. The slope to vertical scaling ratio of f.ff was above the maximum of f.ff
5006-Probe Check Too High Error
5007-Probe Check Too Low Error
5008-Probe Check Delta Too Low Error
5009-Probe Check Delta Too High Error
5010 - Unable to verify positive analyte [analyte name] using curve fitting. X readings were available, but the minimum number of readings required is Y
5011 - Signal Loss detected in the amplification curve
5012 - X probe check failed. The absolute probe check delta value n between reading number m and reading number p was above the maximum of q.
5013 Quantitative value overflow
5014 Quantitative value underflow
5015: Failed to verify valid background slope for analyte [analyte name]. The absolute value of the slope of f.ff was above the maximum of f.ff
5016: Failed to verify valid background error for analyte [analyte name]. The RMS error of f.f was above the maximum of f.f
5017: X probe check failed. Probe check value of n for reading number m was below the valid level of p
5018: Failed to verify valid probe check ratio for analyte [analyte name]. Probe check 1 = m, probe check 2 = n, ratio = f.ff greater than maximum f.ff
5019: Failed to verify valid probe check ratio for analyte [analyte name]. Probe check 1 = m, probe check 2 = n, ratio = f.ff less than minimum f.ff
5020 - Calculations experienced an unexpected error. The final result cannot be calculated.
5021 - Error 5021: Analyte [analyte name] calculated result [calculated value name] is missing. The final result cannot be calculated.
5022 - Error 5022: A calculated value is not a number. The final result cannot be calculated.
5023 - Error 5023: Clinical factor [name] value [value] is not an integer
5024 - Signal loss detected in the growth curve for analyte [<analyteName>]. xx.x decrease in signal with yy.y% decrease at time zz second(s)
5501 - Missing data points for [<data item name>]
5524 - Signal to background ratio value of [S2B value] was below the minimum of [min baseline PCR].
5525 - Mean signal at critical cycle [mean signal value] was below the minimum of [min mean signal value].
5526 - Normalization Probe stability value at last cycle [actual value] was below the minimum of [min value].
5527 - Mean background at critical cycle [mean background value] was above the maximum of [max mean background value].
5536 - [<analyte name>] primer binding check failed. Primer binding check value of <pq0 value> was below the minimum of <min pq0 value>.
5537 - [<analyte name>] primer binding check failed. Primer binding check value of <pq0 value> was above the maximum of <max pq0 value>.
5538 - Analyte [<analyte name>] has an amplitude of <amplitude at critical value> at critical cycle which fails to meet the threshold of <Assay amplitude threshold value>.