## Funktionell Sum
Minor Functional | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error
2008 | NEG_10_0_06+ | 614390505 | Negative Control 1 | 847922 | 701892 | 2025-10-26T23:42:50+0100 | Aborted | ERROR | 92 | Error 2008: Syringe pressure reading of 100.1 PSI exceeds the protocol limit of 100.0 PSI
2037 | B01NOV24B_03_2_19X | 614337470 | Positive Control 2 | 802069 | 635193 | 2025-10-26T20:33:47+0100 | Aborted | NO RESULT | 20 | Error 2037: The cartridge integrity test failed at valve position 1390. The pressure change of -0.6 PSI did not exceed the requirement of 4.0 PSI. The pressure increased from 3.5 PSI to 3.0 PSI during the test
5007 | C15NOV23D_02_1_12X | 614334122 | Positive Control 1 | 839032 | 210058622 | 26/10/25 03:31:57 | Aborted | ERROR | 17.4 | Error 5007: [CT1] probe check failed. Probe check value of 17 for reading number 2 was below the minimum of 110¦Error 5007: [SPC] probe check failed. Probe check value of 19 for reading number 2 was below the minimum of 119
SPC=0 | B01NOV24B_04_2_19X | 614341778 | Positive Control 2 | 802069 | 659869 | 2025-10-26T20:34:11+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 19.7
Delamination | C15NOV23D_07_1_13X_D2B,11,05 | 614349478 | Positive Control 1 | 839032 | 730217 | 2025-10-26T15:01:05+0100 | Aborted | ERROR | 26 | Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 1 was above the maximum of 223¦Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 3 was above the maximum of 223
Indeterminate | C15NOV23D_10_1_15+ | 614389737 | Positive Control 1 | 839032 | 210059038 | 2025-10-26T22:36:21+0100 | Done | MTB DETECTED LOW;RIF Resistance INDETERMINATE | 16.4
Major Functional | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error
False Positive | NEG_03_0_06+ | 614336421 | Negative Control 1 | 847922 | 750444 | 2025-10-26T05:54:16+0100 | Done | MTB Trace DETECTED;RIF Resistance INDETERMINATE | 13.1
False Positive | NEG_03_0_07+ | 614335759 | Negative Control 1 | 847922 | 220000161 | 2025-10-26T05:54:41+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 16.6
Instrument fel | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error
2005 | C15NOV23D_03_1_14X | 614338042 | Positive Control 1 | 839032 | 656159 | 2025-10-26T04:57:00+0100 | Aborted | ERROR | 19.2 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI
1002 | C15NOV23D_03_1_15+ | 614338040 | Positive Control 1 | 839032 | 210059038 | 2025-10-26T04:57:26+0100 | Done | ERROR | 24.6 | Error 1002: The temperature difference of 5.3 ºC exceeds the limit of 5.0 ºC. The temperatures for heaters A and B are 97.3 ºC and 92.0 ºC
1001 | C15NOV23D_03_1_16+ | 614338465 | Positive Control 1 | 839032 | 210012379 | 2025-10-26T04:57:53+0100 | Done | ERROR | 17.3 | Error 1001: The actual temperature 61.6 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.5 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.4 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 61.7 ºC has drifted too far away from the setpoint of 61.0 ºC
2005 (S) | C15NOV23D_06_1_13AX | 614345586 | Positive Control 1 | 839032 | 210018949 | 2025-10-26T13:14:16+0100 | Aborted | ERROR | 18 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI
Ersättningar | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error
I | C15NOV23D_03_1_14AX | 614338044 | Positive Control 1 | 839032 | 656159 | 2025-10-26T08:58:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20
I | C15NOV23D_03_1_15A+ | 614338046 | Positive Control 1 | 839032 | 210058622 | 2025-10-26T08:59:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.6
I | C15NOV23D_03_1_16A+ | 614338467 | Positive Control 1 | 839032 | 652304 | 2025-10-26T09:00:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 15.3
S | C15NOV23D_04_1_14AX | 614341834 | Positive Control 1 | 839032 | 210018949 | 2025-10-26T13:13:00+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.1
I | C15NOV23D_06_1_13AAX | 614345588 | Positive Control 1 | 839032 | 690802 | 2025-10-26T17:29:12+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.9
S | C15NOV23D_06_1_16A+_D06 | 614348455 | Positive Control 1 | 839032 | 210035802 | 2025-10-26T13:15:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 28.4
S | NEG_01_0_08A+ | 614332245 | Negative Control 1 | 847922 | 730163 | 2025-10-26T02:34:57+0100 | Done | MTB NOT DETECTED | 15.5
Observation | Sample ID | Cartridge S/N | Test Type | Instrument S/N | Module S/N | Start Time | Status | Test Result | Max Pressure (PSI) | Error
Delamination | C15NOV23D_08_1_14X_D2B | 614382746 | Positive Control 1 | 839032 | 690802 | 2025-10-26T17:32:06+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 26.6
Delamination | C15NOV23D_08_1_16+_D05 | 614384577 | Positive Control 1 | 839032 | 638845 | 2025-10-26T17:32:55+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.7